## აპარატურა+ინვენტარი
Position | specification | Quantity Required | Specification Offered | Photo | Link | Quantity Offered | Manufacturing country | Manufacturer | Condition | Year of Manufacturing | Model | Unit Price | Total Price | Currency | unit price in gel | total price in gel | Delivery time | Warranty | Comment | EUR | usd
X-RAY შირმა | 1800*900mm | 2
ამბუ სილიკონის ნიღბით მოზრდილთა |  | 15
ამბუ სილიკონის ნიღბით ნეონატალური |  | 5
ამბუ სილიკონის ნიღბით პედიატრიული |  | 10
ანესთეზიის აპარატი | - Anesthetic system for anesthesia for adults, children and neonatal, with the ability to work in a half-open and half-closed circuit and low-flow anesthesia. Color touchscreen display not less than 10 inches; with flowmeters; Built-in ventilator with pneumatic drive. Modes of Ventilation: Volume control; Pressure control; CPAP/BIPAP; With the possibility of mechanical breathing; with oxygen and flow sensor; The main unit of the apparatus with inlets for minimum two gases (O2, Air); O2 and Air Pressure range at inlet: 3-6 bar or better; with DIN standart connector; on a mobile base (with 4 wheels; with a brake;) with minimum one drawers. Minimum two Anesthetic vaporizer (sevoflurane; isoflurane; with adapter for filling liquid, if it's necessary). - with patient monitoring system - SpO2; NIBP; IBP; TEMP; CO2; sidestream multigas analyzer for anesthesia (A multigas analyzer can be integrated on the anesthesia machine) - with pipelines for O2 and Air (DIN standard) - 2GAS FLOWMETER, O2 AIR, - Built-in battery with up to 30 minute of operation - BAG ARM SUPPORT - AGSS - Absorber canister assembly, reusable - The device should be ready for use - AC power input: 220V 50Hz - Warranty at least 12 month - Certificates CE or FDA | 1
ანესთეზიის აპარატი | - Anesthetic system for anesthesia for adults, children and neonatal, with the ability to work in a half-open and half-closed circuit and low-flow anesthesia. Color touchscreen display not less than 10 inches, electronic flowmeters with graphic display. Built-in ventilator with pneumatic drive. Modes of Ventilation: Volume control; Pressure control; CPAP/BIPAP; With the possibility of mechanical breathing; with oxygen and flow sensor; The main unit of the apparatus with inlets for minimum two gases (O2, Air); O2 and Air Pressure range at inlet: 3-6 bar or better; with DIN standart connector; on a mobile base (with 4 wheels; with a brake;) with minimum one drawers. Minimum two Anesthetic vaporizer (sevoflurane; isoflurane; with adapter for filling liquid, if it's necessary). - with patient monitoring system - SpO2; NIBP; IBP; TEMP; CO2; sidestream multigas analyzer for anesthesia (A multigas analyzer can be integrated on the anesthesia machine) - with pipelines for O2 and Air (DIN standard) - 2GAS FLOWMETER, O2 AIR, - Built-in battery with up to 30 minute of operation - BAG ARM SUPPORT - AGSS - Absorber canister assembly, reusable - The device should be ready for use - AC power input: 220V 50Hz - Warranty at least 12 month - Certificates CE or FDA | 1
ბავშვთა გასასინჯი სპეციალური მაგიდა |  | 1
გამათბობელი ლეიბი |  | 1
გამათბობელი ჰაერის დაბერვით |  | 1
გამაციფრულებლის კასეტა დიდი | Cassette for computed radiography system (CR), must be compatible for Fujifilm FCR PRIMA T; Size 35.4x43.0 cm; | 2
გამაციფრულებლის კასეტა პატარა | Cassette for computed radiography system (CR), must be compatible for Fujifilm FCR PRIMA T; Size 24x30 cm; | 2
გორგოლაჭებიანი სკამი | უჟანგავი ლითონის (მრგვალი ზედაპირის) | 1
დაბალი წნევის ამომქაჩი | Suction Aspirator; With jar holder; Jar size 1 litre with lid, with overflow valve system; Antibacterial and hydrophobic filter at least 3 pieces; Flow rate 20 l/min or better; Suction pressure 580 mm/Hg or better (adjustable); Noise level: less than 61 dBA; The device should be ready for use; AC power supply: 220V 50Hz; warranty at least 1 year; Certified with CE or FDA ; | 10
დეფიბრილატორი | Biphasic defibrillator; At least 7 inc LCD display; Energy range 10-360 J or better (Adjustable, Biphasic Waveform Technology); 360J charging time not more than 10 second; Built-in standard 12-lead ECG with ECG cabble; with Synchronous and asynchronous mode; With Semi-automatic (AED), manual and paicing modes, With AED pads (at least 3 pairs) and cabble for AED pad connection; With paddle type electrode (Apex, Sternum) for adult and pediatric patients; defibrillation operation should be possible manage from paddles; with thermal printer; Built-in battery with up to 4 hour of operation; The device should be ready for use; AC power supply: 220V 50Hz; warranty at least 1 year; Certified with CE or FDA; | 10
დიალიზის სავარძელი |  | 10
დიფერენციალური წნევის მანომეტრი |  | 22
ელექტროკარდიოგრაფი 12 განხრიანი | 12 channel electrocardiograph; with 12 lead ECG cable, with full Set Chest Suction Ball & Limb Clamp Clip ECG Electrodes for adult; with integrated thermal printer; Thermal paper size A4; At least 7 inc color display (touch screen); with internal storage for patient database; Should include an interpretive function; with battery; at least 2 hours battery life; With QWERTY keyboard for entering patient data; should be possible to preview the full print directly on the screen; The device should be ready for use; AC power supply: 220V 50Hz; warranty at least 1 year; Certified with CE or FDA ; | 10
ელექტროკარდიოგრაფის კაბელები |  | 50
ელექტროკარდიოგრაფის კიდურების და გულმკერდის სამაგრები |  | 100
ენდოსკოპის ელექტრო სარეცხი მანქანა | ენდოსკოპის გამრეცხი მანქანა | 2
ენდოსკოპიური სისტემა (ნეონატალური, პედიატრიული, მოზრდილთა) მეორადი | ერთი გასტრო; ერთი კოლონო; ერთი ბრონქოსკოპი სისტემაზე და 1 ვიდეო ბრონქოსკოპი; | 1
ვაკუუმის რეგულატორი |  | 3
ვაკუუმის რეგულატორი ვენტური ტიპის |  | 9
ვაკუუმის ქილა |  | 9
თირკმლისებული თასი პლასტმასის |  | 30
ინფუზომატი ორშპრიციანი | Dual-channel Syringe Pump; With not less than 3 inches LCD display; It should be possible to operate two syringes together and separately; With fast infusion button; Compatible with 10, 20, 30, 50, 60 ml syringes; With the ability to automatically recognize the size of the syringe; Flow rate range 1-1000 ml/h or better; Minimum increment not more than 1 ml/h; Flow rate accuracy ±2% or better; With battery; Battery life is not less than 4 hours; Waterproof design; With tripod holder; Audible & Visual alert (empty, occlusion, low battery, AC power fail, end of infusion, etc.); The device should be ready for use; AC power supply: 220V 50Hz; warranty at least 1 year; Certified with CE or FDA ; | 35
ინფუზომატი სისტემის (პარენტერალური) | Infusion pump; With not less than 3.5 inches LCD display; A standard catheter may be used (not just the manufacturer's chip catheter); With fast infusion button; Minimum infusion rate 0.1 ml/h or better; Maximum infusion rate 1400 ml/h or better; Accuracy ±5% or better; With air bubble detection; With tripod holder; With battery; Battery life is not less than 4 hours; With audible & Visual alert; The device should be ready for use; AC power supply: 220V 50Hz; warranty at least 1 year; Certified with CE or FDA ; | 17
ინფუზომატი სისტემის (ქიმიოთერაპიისთვის) BBRAUN |  | 3
კარდიოტოკოგრაფი |  | 1
კერატორეფრაქტომეტრი |  | 2
კოლპოსკოპი |  | 1
ლაბორატორიული მიკროსკოპი |  | 3
ლარინგოსკოპი (ვიდეო) | Video Laryngoscope, with reusable blades; At least 3 MAC type blades (size: 1,2,3); Display size not less than 3 inc; With 2.0 mega-pixel camera or better; with button for take picture; battery continuous working time more than 150 Min; with charger (AC power supply: 220V 50Hz); Illuminance: ≥ 700 lux; The device should be ready for use; warranty at least 1 year; Certified with CE or FDA; | 3
ლარინგოსკოპი ოპტიკური შუქგამტარით | Fiber optic Laryngoscope; Bulb must be in handle; with at least one spare bulb; light intensity not less than 800 lux; with at least 5 pcs fiber optic blade, type of MAC (sizes: 0,1,2,3,4); blades must be autoclavable at 134 degrees celsius; possibility to use disposable battery (specify battery type); warranty not less than 6 month; Certified with CE or FDA; | 6
ლარინგოსკოპი რთული ინტუბაციის |  | 1
ლეიკოციტების მთვლელი C5 |  | 2
მაღალი წნევის ამომქაჩი | Suction aspirator; With two 2 or 2.5 liter Jars with lid and overflow valve system; With jars holder; Additionally Antibacterial and hydrophobic filter at least 5 pieces; Flow rate 60 l/min or better; Suction pressure 670 mmHg or better (adjustable); noise level less than 55 dBA; Pump must be Oiless and maintenance free piston type; The device should be ready for use; AC power supply: 220V 50Hz; warranty at least 1 year; Certified with CE or FDA ; | 4
მედიკამენტების კარადა ორკარიანი |  | 32
მეიოს მაგიდა | size:600*400*900-1500mm Height adjustable | 4
მონიტორი CO2 | Patient monitor; With colod LCD display at least 12 inc; Parameters: SpO2, Pr, Nibp, 2 Ibp, Etco2, ECG, 2-Temp; Suitable for Adult and Pediatric patients; Accessories: 1 pcs SpO2, 1 pcs 5 lead ECG cable, 2 pcs Cuff (Large, Medium) and 1 pcs tube, 2 pcs plug for ibp1 and ibp2, 2 pcs catheter for ibp, 1 pcs cup for Etco2, 1 pcs tube for Etco2, 1 pcs skin temp probe; With audible & Visual alert (Alarm sound level must be adjustable); With battery; Battery life is not less than 2 hours; The device should be ready for use; AC power supply: 220V 50Hz; warranty at least 1 year; Certified with CE or FDA; | 8
მონიტორი ინვაზიური | Patient monitor; With colod LCD display at least 12 inc; Parameters: SpO2, Pr, Nibp, 2 Ibp, ECG, 2-Temp; Suitable for Adult and Pediatric patients; Accessories: 1 pcs SpO2, 1 pcs 5 lead ECG cable, 2 pcs Cuff (Large, Medium) and 1 pcs tube, 2 pcs plug for ibp1 and ibp2, 2 pcs catheter for ibp, 1 pcs skin temp probe; With audible & Visual alert (Alarm sound level must be adjustable); With battery; Battery life is not less than 2 hours; The device should be ready for use; AC power supply: 220V 50Hz; warranty at least 1 year; Certified with CE or FDA; | 1
მონიტორი საბაზისო | Patient monitor; With colod LCD display at least 12 inc; Parameters: SpO2, Pr, Nibp, ECG, 2-Temp; Suitable for Adult and Pediatric patients; Accessories: 1 pcs SpO2, 1 pcs 5 lead ECG cable, 2 pcs Cuff (Large, Medium) and 1 pcs tube, 1 pcs skin temp probe; With audible & Visual alert (Alarm sound level must be adjustable); With battery; Battery life is not less than 2 hours; The device should be ready for use; AC power supply: 220V 50Hz; warranty at least 1 year; Certified with CE or FDA; | 58
მონიტორი საბაზისო ნეონატალური | Patient monitor; With colod LCD display at least 12 inc; Parameters: SpO2, Pr, Nibp, ECG, 2-Temp; Suitable for Adult, Pediatric end Neonatal patients (It should be possible to select from software); Accessories: 1 pcs SpO2 for adult, 1pcs SpO2 for neonatal (reusable), 2 pcs SpO2 for neonatal (disposable), 1 pcs 5 lead ECG cable, 3 pcs Cuff (Adult, Child, Neonatal) and 1 pcs tube, 1 pcs skin temp probe; With audible & Visual alert (Alarm sound level must be adjustable); With battery; Battery life is not less than 2 hours; The device should be ready for use; AC power supply: 220V 50Hz; warranty at least 1 year; Certified with CE or FDA; | 1
მონიტორი საბაზისო პედიატრიული, მოზრდილთა | Patient monitor; With colod LCD display at least 12 inc; Parameters: SpO2, Pr, Nibp, ECG, 2-Temp; Suitable for Adult, Pediatric end Neonatal patients (It should be possible to select from software); Accessories: 1 pcs SpO2 for adult, 1pcs SpO2 for neonatal (reusable), 1 pcs 5 lead ECG cable, 3 pcs Cuff (Adult, Child, Neonatal) and 1 pcs tube, 1 pcs skin temp probe; With audible & Visual alert (Alarm sound level must be adjustable); With battery; Battery life is not less than 2 hours; The device should be ready for use; AC power supply: 220V 50Hz; warranty at least 1 year; Certified with CE or FDA; | 10
მონიტორის ეკგ გადამწოდი |  | 100
მონიტორის სატურაციის გადამწოდი |  | 200
მონიტორის წნევის მანჟეტი |  | 200
მონიტორის წნევის მანჟეტის მილი |  | 50
ნაწოლის საწინააღმდეგო ლეიბი |  | 5
ორდონიანი საპროცედურო მაგიდა უჟანგავი | სამანიპულაციო მაგიდა | 7
ოტოსკოპი |  | 3
ოფთალმოსკოპი |  | 5
პანელი | რეანიმაციის | 9
პანელი | საპალატე | 11
პაცენტის მონიტორის სადგამი | კედლის კალათით | 27
პორტი | ჟანგბადის | 100
პორტი | co2 | 10
პორტი | აირის | 60
პორტი | ვაკუუმის | 30
პორტი | ვენტურის | 9
რადიაციული დამცავი გონადები ზომებით |  | 4
რადიაციული დამცავი ჟილეტი ქვედაბოლოთი |  | 6
რადიაციული დამცავი სათვალე |  | 2
რადიაციული დამცავი საყელო (მოზრდილთა და პედიატრიული) |  | 7
სამანიპულაციო ურიკა | ორდონიანი , უჟანგავი ფოლადი | 14
სანარკოზე აპარატის CHIRANA VENAR გაზანალიზატორის რეზერვუარი (ჭიქა) |  | 3
სანარკოზე აპარატის მინდრეის WATO EX-55 PRO გაზანალიზატორის რეზერვუარი (ჭიქა) |  | 3
საოპერაციო მოძრავი სკამი |  | 1
სასწორი ახალშობილის |  | 2
სასწორი დიალიზისთვის |  | 1
საშრობი კარადა ორ განყოფილებიანი უჟანგავი მეტალის | ენდოსკოპიის საშრობი კარადა | 1
საწოლი | ფუქციური მექანიკური | 10
საწოლი | CPR ფუნქციით, მექანიკური 4 ფუნქციით ან ელექტრო | 26
სიმაღლის საზომი |  | 10
სპირომეტრი |  | 10
სტერილიზატორის ბიქსი | ოთკუთხედი | 3
სტერილიზატორის ბიქსი ოთკუთხედი |  | 3
სტრეტჩერი | 1.8მ სიგრძის | 1
სტრეტჩერი |  | 9
ტრედმილი |  | 1
ტუმბო | ABS მატერიალი ლურჯი ტუმბო 2 უჯრიანი, წმენდადი ზედაპირით | 73
ფეტალური დოპლერი |  | 2
შესახვევი მაგიდა უჟანგავი ფოლადის ორდონიანი | სტერილური იარაღების გასაშლელი მაგიდა, უჟანგავი ფოლადის | 4
შირმა | 3 ფრთიანი | 7
შოკის მაგიდა | ABS მატერიალი | 21
შტატივი | ოთხსაკიდიანი, ოთხი ბორბლით | 9
ცენტრიფუგა |  | 2
წნევის საზომი აპარატი | მექანიკური | 1
ჰოლტერი |  | 7

## Sheet2 (hidden)
Condition'
New
Used
Demo
Refurbished