## აეკი
დასახელება | სპეციფიკაცია | რაოდენობა
აეკი (რეფაბრიკირებული/მეორადი) | 1. Stockert S5 - Console (movable) on wheels for 5 pumps - four stands + one had pump mounted on a mast. 2. System control panel - integrated control for the entire systems 3. 4 Peristaltic pumps (roller pumps) - single-head, one (had pump) mounted on a mast. 4. Double-head peristaltic pump – for cardioplegia 5. Hemodynamic control system - monitoring modules 5.1. Micro air embolism and liquid level detectors 5.2. temperature measurement module 5.3. Timing module - triple stopwatch, 5.4. pressure measurement module - minimum two-channels, 5.5. Cardioplegia module 6. Air-Oxygen gas blender – for example Sechrist or electronic (more expensive) 7. Vacuum assisted venous drainage controller – like this(https://www.perfusion.com/product/digital-vavd-vacuum-regulator-red-0-100mmhg/) or minimum manual venous line occluder 8. Electrical Venous Occluder (EVO) (https://www.livanova.com/en-GB/Home/Cardiopulmonary/Heart-Lung-Equipment/EVO.aspx) 10.Optional accessories 10.1.perfusionist chair 10.2. LED lamp | 1 კომპლექტი

## ულტრაბგერითი ვაკუუმ საქაჩი
დასახელება | სპეციფიკაცია |  | რაოდენობა
ულტრაბგერითი ვაკუუმ საქაჩი (ახალი) | ნეიროქირურგიული და ღვიძლის ქირურგიის ულტრაბგერითი ასპირაციის გენერატორი |  | 1 კომპლექტი
მართვის პედალი
სადგამი ურიკა
| სამუშაო ინსტრუმენტი (Handpiece) ნეიროქირურგიისათვის | 36 kHz
| სამუშაო ინსტრუმენტი (Handpiece) ღვიძლზე სამუშაოდ | 23 kHz
| თავაკი (36 kHz Handpiece-სთვის) | არანაკლებ 5 ცალი განსხვავებული ზომის
| თავაკი (23 kHz Handpiece-სთვის) | არანაკლებ 4 ცალი განსხვავებული ზომის
| სამუშაო მილების ნაკრები | არანაკლებ 10 ნაკრები
| თავაკების საფიქსაციო მოწყობილობა (36 kHz Handpiece-სთვის) | მრავალჯერადი (ავტოკლავირებადი)
| თავაკების საფიქსაციო მოწყობილობა (23 kHz Handpiece-სთვის) | მრავალჯერადი (ავტოკლავირებადი)
| სასტერილიზაციო კონტეინერი | თავაკებისთვის და სამუშაო ინსტრუმენტებისათვის (მრავალჯერადი, ავტოკლავირებადი)
თვითტესტირების რეჟიმი
ძვალზე მუშაობის შესაძლებლობა
| მონიტორი | სენსორული ეკრანი (Touch Screen)

## ნეირო მიკროსკოპები
დასახელება | სპეციფიკაცია |  | რაოდენობა
NeuroSurgical Microscope (new) | With three binocular module | a surgeon tube, lateral and opposite stereo assistant tube a binocular tube 30°-150° variable a binocular tube 30°-150° variable a binocular tube 0° - 180° variable | 2 კომპლექტი
| Optics | 540° rotation 50° lateral tilt to left and right -30° /+120° inclination tilt
| Modularity | -oncological fluorescence observation filter module -vascular fluorescence
| Monitor | 27" Medical, minimum HD, Touch screen
| Camera | Integrated minimum HD
| Documentation systems | 4K
| FusionOptics | For increased depth of field and high resolution for main surgeon and opposite assistant
| Magnification | 6:1 zoom, motorized
| Total magnification | 1.0× to 12.1× with 10× eyepiece
| Focus | Motorized via multifocal lens, with manual adjustment
| Objective / working distance | 225–600 mm, motorized multifocal lens, continuously adjustable and manual adjustment option
| Field of view | 17.4 to 210 mm ø with 10× eyepiece
| Eyepieces | Wide-field eyepieces for persons wearing glasses 8.3×, 10× and 12.5× dioptric adjustment, ±5 diopter settings and adjustable eyecup
| Integrated 360° rotatable adapter | For main surgeon binocular (IVA, ULT) and opposite assistant (ULT)
| Illumination | - High-output 2x 400-W redundant xenon arc-lamp systems via fiber optics cable
| Intraoperative balancing | Automatic intraoperative AC/BC balancing of AC and BC axes
| Brakes | Floor stand
| Control unit | - Programmable touch-screen with userfriendly Graphical User Interface - Built-in electronic auto-diagnosis and user support - Software independent hard keys for illumination and auto-balancing - Indicator for main/backup illumination and fluorescence modes
| Control elements | -Pistol handle with programmable Functions -Optional mouthswitch -Optional wireless Footswitch
| Carrier for monitor | flexible arm with 4 axis for rotation and inclination
| Construction | 360° rotating Casters
| Power connection | 220 V

## ნეირო მონიტორინგი (პორტატული)
დასახელება | სპეციფიკაცია | რაოდენობა
ნეირო მონიტორინგი | პორტატული | 1 კომპლექტი

## სტერეოტაქსია+სტიმულაცია
დასახელება | სპეციფიკაცია | რაოდენობა
Stereotaxy headring unit CT/MR compatible for convergence radiation | Titanium headring, 4 universal head fixing post artefact free with each 4 fixing pins 40 mm, 50 mm 60 mm and 70 mm | 1
| Headring, titanium | 1
| Universal head fixing post, free of artefacts, titanium | 4
| Opti pin fixing pin 40 mm | 1
| Opti pin fixing pin 50 mm | 1
| Opti pin fixing pin 60 mm | 1
| Opti pin fixing pin 70 mm | 1
| Headring holder for Mayfield/Kees | 1
| Socket wrench | 2
| Torque wrench 70Ncm | 1
| Universal CT holder, 280 x 300 x 75 mm | 1
| CT/MR Localisation set for Titanium & Aluminium Headring | 1
| RM Stereotaxy unit: RM aiming bow with instrument carrier, depth control, depth control stop, vertical arc 105° and 85°, vertical arc support, probe holder, dovetail guide, lens NV and NS, 2 U-shaped base, sterilisation box and instructions for use. | 1
| RM biopsy probe 2.5mm, work piece length 280mm | 3
| RM test needle | 1
| Target point simulator. Adjusting unit for X-Y-Z axis. | 1
| Headring, aluminium | 1
| Target point simulator | 1
| Sterilisation box or target point simulator | 1
| Positioning Aid for stereotactic Systems | 1
| MicroDrive test mandrin for phantom target point for ZD/RM, Use only with MicroDrive and Microgun. OD = 1.65 mm, effective length 200 mm, packed individually, autoclavable | 1
| PS-Planning System | 1
| Planning Software PS, 3D stereotactic applications, including DICOM image import module | 1
| Fusion module for PS automatic coregistration | 1
| PC for Planning Software PS, with perating system - 1 x 24" or higher wide screen full HD or 4k - DVD Drive / burner - Colour laser printer | 1
| PS Copy protection dongle | 1