**Micro-ohmmeter**

Our company needs Micro-ohmmeter for testing following devices:

1. Vacuum circuit breaker
2. Power transformer up to 2500 kVA power, 10 kV voltage
3. Squirrel cage motor up to 800 kW power, 6 kV voltage
4. Current and voltage transformers up to 10 kV voltage

Technical characteristics:

1. Measurement range - 1µΩ up to 2500 Ω
2. Measurement current – 1mA up to 10 A
3. Accuracy class – 0.05%
4. Communication output – RS232
5. Measurement type – 4 wire kelvin test
6. Power supply – rechargeable Battery

Accessories:

1. Test leads - 10A kelvin clamps, length ≥3 m – 1 set.
2. Caring bag – 1 pc.
3. User’s manual on English/Russian language – 1 set
4. Data transfer software for PC – 1 pc.
5. Communication cable RS232 – 1pc.
6. Source cable – 1 pc.